### BCA Protein Assay Protocol

- Prepare diluted albumin (BSA) standards according to the table below. Use the same diluent as the samples.

| Vial | Volume of diluent (uL) | Volume and source of BSA (uL) | Final BSA Concentration (ug/mL) |
| --- | --- | --- | --- |
| A | 0 | 300 of stock | 2000 |
| B | 125 | 375 of stock | 1500 |
| C | 325 | 325 of stock | 1000 |
| D | 175 | 175 of vial B | 750 |
| E | 325 | 325 of vial C | 500 |
| F | 325 | 325 of vial E | 250 |
| G | 325 | 325 of vial F | 125 |
| H | 400 | 100 of vial G | 25 |
| I | 400 | 0 | 0=blank |

- Calculate how much working reagent is required using this formula:

(#standards + #unknowns) x (#replicates) x (200 uL) = volume WR needed

- Prepare working reagent by mixing 50 parts of BCA reagent A with 1 part of BVA reagent B.

Ex. For 50 mL of WR needed, combine 50 mL of A with 1 mL of B.

- Pipette 25 uL of each standard or unknown sample replicate into a microplate well.

- Add 200 uL of the working reagent to each well (using multichannel) and mix thoroughly on a plate shaker for 30 seconds.

- Cover plate and incubate at 37C for 30 minutes.

- Cool plate to RT. Measure the absorbance at or near 562 nm on a plate reader.

- Subtract the average absorbance of the blank standard from all other standard and unknown sample replicates.

- Prepare a standard curve by plotting the average blank-corrected measurement for each BSA standard vs. its concentration in ug/mL. Use the standard curve to determine the protein concentration of each unknown sample.

| A1 | A2 | A3 | I1 | I2 | I3 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| B1 | B2 | B3 |  |  |  |  |  |  |  |  |  |
| C1 | C2 | C3 | Sample 1 | Sample 2 | Sample 3 |  |  |  |  |  |  |
| D1 | D2 | D3 |  |  |  |  |  |  |  |  |  |
| E1 | E2 | E3 |  |  |  |  |  |  |  |  |  |
| F1 | F2 | F3 |  |  |  |  |  |  |  |  |  |
| G1 | G2 | G3 |  |  |  |  |  |  |  |  |  |
| H1 | H2 | H3 |  |  |  |  |  |  |  |  |  |