- Setup Videos

- Software

  - Triax Software

- Relevant Documents

  - Triax manual

  - Gradient Master manual

  - Piston Gradient Fractionator manual

  - Supplemental Gradient Station Manual

  - Gradient Primer

- Non-Supplied Items to gather

  - Tools

    - Wrench or pliers to remove the bolts in the wooden shipping plate

    - Razorblade

    - Fine forceps to remove tubes from the rotor buckets

  - Lab Gear

    - Cold rotor and centrifuge, both reserved for the day

      - Big Screwdriver for removing the lids from the buckets

      - Spinkote lubricant if the o-rings on the bucket are needing rejuvenation

    - Sample on ice that has been spun for at least 1 min at 14K.

    - Heavy and light vacuum filtered sucrose solutions, at room temperature (use something like the Millipore SCGP00525 to filter them).

    - 20mL of 70% isopropyl alcohol

    - A few mL of Trypan Blue

    - Filtered rinse water (1 liter). MiliQ is OK

    - 30-60 cc syringe for layering the light solution

    - 10 cc syringe for layering the heavy solution

    - A 20 or 30 cc syringe filled with MilliQ water connected to the rinse adaptor for the zeroing of the Triax

    - Rack for Centrifuge tubes

    - Parts to make Sample layering device

      - 1 cc syringe per trainee

      - 1 cc graduated plastic pipette per trainee

      - 1 piece rubber tubing to connect the above per trainee

    - Large orifice dH2O squirt bottle

    - 100 ml beaker for the rinse pump purge on right side of Station

    - Large heavy bottle for the Gilson gutter overflow

    - 10-milligram scale (0.00 g)

    - Fridge or cold room at 4°C

- BioComp supplied Equipment

  - Piston Gradient Fractionator and Gradient Master/ Gradient Station

      - Kit

        - Cannula for both the layering syringes

        - Vial level

        - Philips screwdriver

        - Fine tip marker

        - Rinse adaptor

        - Silicon lubricant

        - Tube Conditioning Bolt

        - 60cc syringe for air dry

        - 60 cc syringe without plunger for rinse reservoir

      - Gradient Forming Tube Holder

        - Tube caps

        - Marker Block

      - Fractionating Tube Mount

        - Tip and Seal

        - Triax Flow Cell

      - Seton Centrifuge tubes

    - Triax Flow Cell

      - Idex Fitting with blue ferrule

    - Gilson Fraction Collector

      - Fraction Collector to Flow Cell tubing (290mm)

      - Tube Rack/ 96 well rack

      - Eppendorfs as needed

    - Laptop with Triax software installed

    - IF install is via video call -  A second laptop with good, fully charged battery and SKYPE installed, located so the camera has a good view of where the system will be installed.

Agenda

- Setting up the instruments

  - To save time you can use these videos to get the instruments set up ahead of time.  Otherwise, we will guide you through the process.

- Making a Gradient

  - Video Overview

  - Overview of functionality including:

    - Levelling Plate

    - Programming a NEW run

    - LAST and RCNT runs

    - Finding a run in the LIST

  - Lubricating the tubes.  Please review this video

  - Layering the heavy and light solutions.

  - If possible make one extra Gradient with Trypan Blue added to the Heavy so we can scan the gradient alone and see if it is linear.

  - Make the Gradients

  - Put the Gradients in the cold room for 45 minutes.

- Prepare for Sample layering

  - Make Sample Layering Devices (Video)

  - Ensure the scale is well located and tares reliably

  - Fill a large bucket with ice.

- Sample Layering

  - Please review this document

  - Tare tubes

  - Layer Samples

- Load Gradients into a centrifuge

- 1 hr break for lunch

- Fractionation, Profiling and Fraction Collection

  - Review Software

    - INTRO

    - USER

    - LED POWER

    - SCAN SETUP

    - GRAPH

    - Zeroing

  - Profile Trypan Blue Gradient

    - Ensure everything is functioning normally including Gilson.

- Profile Gradients

  - Use OVERLAY GRAPH and COMPARE DATASETS

- Clean Instrument

  - Review this document

  - Video