# Plate Spinner and Lightcycler for qrtPCR

Using plate spinner in INBRE Core Lab

- Ensure proper rotor is in place

- Place 96-well plate into plate holder and place into centrifuge

- Gently close lid and spin at 500xg for 2 minutes

Using Roche LightCycler

- Turn on machine and wait for the left light on the front to turn green. Once it has, press and hold the  button on the front of the machine for a moment, to open the 96-well plate tray

- Place plate into tray, aligning the slanted edge correctly

- Open the LightCycler application and long in with the following credentials:

Username: K_Ramsey Lab

Password: Roche480

- Select New Experiment from template > Run template > ‘Kramsey_lab_old_stepone’ and press the check button

- Ensure the program matches the program specs listed in ‘Real-time PCR programs.docx’

- Select ‘Start Run,’ name and save the experiment to the KRamsey Lab > Experiments folder

- The program will take ~1.5 hours to complete

- To export:

  - Select Analysis, and click the ‘+’ button. The Analysis Type is ‘Absolute Quant/2nd Derivative Max’

    - Select appropriate subsets then hit the check

  - Select ‘Calculate’

  - Scroll to the right in the samples column, selecting ‘CP’ for Crossing Point then save the Analysis with the save icon

  - Select the Compass icon > ‘Results Batch Export’

  - Select the experiment you saved in the KRamsey Lab > Experiment folder, press the arrow to select the experiment then click ‘Next’

  - Browse Target and save to the flash drive

  - Select the same analysis type, select ‘Next,’ then select ‘Start’

  - Ensure it says 1 file exported, if file is not exported, likely you forgot to save on the analysis page.