## Slide 1
Compressed Gas and Cryogen Seminar with a focus on EH&S
October 12, 2022
Host: Nic Englehart - EH&S
Room 234 - Pastore Hall
51 Lower College Road, Kingston RI
Speaker: Ralph Nigro – Airgas Specialty Gas Product Manager

## Slide 2
Safety moment.
Personal Protective Equipment
Phase Separator
Notes: Cryogenic Transfer Hose
Phase Separator – reduces “splashing”
Gloves approved for handling cryogens – not your “garden” variety gloves
Faceshield approved for use with cryogens
Duck? Make yourself like a duck…water flows off duck’s back, that’s what you want to happen when/if a cryogen touches you…if flows right off

## Slide 3
To Tape or Not to Tape?
That is the question?

## Slide 4
Where & where not to use Teflon tape
This is the seating surface
What is wrong with this picture?

## Slide 5: How Do You Select A Regulator?
Match regulator to the application
Product compatibility
  Pressure and chemical
Delivery pressure
Flow
Accuracy
Pressure Regulators
Or Call your account manager.

## Slide 6
The primary purpose of a regulator is to reduce high-pressure gas found in a cylinder to a lower, useable pressure.
Pressure Regulators
2/10
Purpose of a Regulator

## Slide 7: Regulators regulate pressure not flow
Pressure Regulators
Regulator for Pressure
Flowmeter for Flow

## Slide 8: Flowmeters
Must be vertical & not on an angle

## Slide 9: What’s The Difference Between Single and 2-Stage Regulators?
“One Ear”
Single Stage
“Two Ears”
Two Stage

## Slide 10: What’s The Difference Between Single and 2-Stage Regulators?
“One Ear”
Single Stage
“Two Ears”
Two Stage

## Slide 11: What’s The Difference Between Single and 2-Stage Regulators? 1 PSI to 100 PSI 0.1 PSI to 100 PSI
Single Stage
Not Precise
Delivery Pressure
Due to Supply Inlet Effect
Two Stage
Precise Delivery Pressure

## Slide 12: Regulator Troubleshooting Call your account manager

## Slide 13: Flexible Pigtails
Never use flexible pigtails with Teflon inner core
  Contaminates are absorbed when pigtail has no pressure, then drawn into the gas stream when high-pressure gas flows
Molecules (Helium/Hydrogen)are small enough to diffuse out of the hose are small enough to enter into the gas stream when flowing
Always use stainless steel flexible core

## Slide 15: Excess Flow Valves
Shuts off gas during high flow or if a line break occurs.
  Typically sized at 4 to 10x normal flow.

## Slide 16
Gas Cabinets

## Slide 17: Do not identify cylinders by their color Read the label before using any cylinder.
Safety moment

## Slide 18: Cylinder Types
Safety moment - Make sure they are secured
Lecture bottles might seem convenient but the disposal cost could be upwards of $10,000.00

## Slide 19: Handling Cylinders
Cylinders are not Door stops or Lab Coat hangers
Cylinders should never be Dragged.
Always make sure a cylinder cap is on tight before moving.
NEVER lift a cylinder by the Cap.

## Slide 20
What is the proper way to change a regulator?
Regulators, lets get started

## Slide 21
Close the valve clockwise

## Slide 22: Drain the regulator
All gauges should read zero

## Slide 23: Back out the adjusting screw
No pressure or flow is your goal

## Slide 24: Disconnect the regulator
Use a proper wrench

## Slide 25: Preparing to connect the regulator
Move the cylinder into position
Rotate the cylinder so that when you connect the regulator you will be standing to the side and not in front or behind the regulator
Secure it to a bracket with either a chain or strap
  Never stand directly in front of the regulator when it is being pressurized, in the event of a failure serious injury may occur
Remember:
Do not stand in front of the regulator
Do not stand behind the regulator
Do not stand on the regulator side
WHY?

## Slide 26: Hand tight the regulator, then Tighten with a wrench

## Slide 27: Lets open the cylinder valve Barely crack the cylinder valve
Pressure should build up slowly.
Then adjust the regulator to the desired pressure

## Slide 28: Cryogenic Containers

## Slide 29
What is a cryogenic liquid container?
A pressurized, double-walled, insulated
container used to hold either cryogenic
liquefied gas.
The inner vessel is insulated from the outer
vessel by a vacuum space.

## Slide 30
THIS IS THE INCORRECT WAY OF MOVING A CONTAINER REMEMBER ALWAYS PUSH DO NOT PULL

## Slide 31
This is the correct way to move a liquid container

## Slide 32
Common Configurations
Liquid Level Gauge (10)
Liquid and Fill
Valve (2)
Relief Valve (7)
Vent Valve (4)
Rupture Disk (5)
Liquid Valve (2)
Liquid Level Gauge (10)
Pressure Building Valve
(9)
Gas Use Valve (11)
Rupture Disk (5)
Vent Valve
(4)
Relief Valve (7)
Low Pressure Design
High Pressure Design

## Slide 33
Vapor Space
Liquid Product
Inner Vessel
Outer Vessel
Container Cross Section
Vacuum Space
Liquid Level

## Slide 34
Gas Withdrawal
1 Gas withdrawal vaporizer
2 Liquid withdrawal valve
3 Economizer regulator
4 Vent valve
5 Inner tank rupture disk
6 Pressure gauge
7 Pressure relief valve
8 Pressure building regulator
9 Pressure building valve
10 Liquid level gauge
11 Gas withdrawal valve and integral check valve
12 Check valve in house line
13 Outer tank rupture disk
14 Pressure building vaporizer
2
1
3
4
5
6
7
8
9
10
11
12
13
14
Notes: .

## Slide 35
Pressure Building
1 Gas withdrawal vaporizer
2 Liquid withdrawal valve
3 Economizer regulator
4 Vent valve
5 Inner tank rupture disk
6 Pressure gauge
7 Pressure relief valve
8 Pressure building regulator
9 Pressure building valve
10 Liquid level gauge
11 Gas withdrawal valve and integral check valve
12 Check valve in house line
13 Outer tank rupture disk
14 Pressure building vaporizer
2
1
3
4
5
6
7
8
9
10
11
12
13
14
Notes: Dependable product withdrawal is dependent upon a stable cylinder pressure to provide that flow. The previous slide on the economizer explained how the cylinder pressure is limited on the high end. The pressure build system controls the pressure on the low end. The minimum product pressure is controlled by the pressure build system. The pressure build circuit requires the user to open the pressure build valve to activate the pressure build system. The product flow can be traced by color, blue for liquid and green for gas. If the cylinder pressure is below the set point of the pressure build regulator, usually ~125 psig, the regulator opens allowing liquid product into the pressure build vaporizer. In the vaporizer the liquid is converted to a gas which travels through the pressure build regulator into the vapor space of the cylinder. When the pressure in the vapor space builds to satisfy the regulator set point, the regulator closes.

## Slide 36: Dewar filling, a few points
PPE:
Safety Glasses
Face Shield
Cryogenic Gloves
Long Sleeves
Pants without a cuff
Phase Separator

## Slide 37: Liquid nitrogen expansion ratio is 696 to 1

## Slide 38: Wrenches
Do not use pliers. Use an appropriate cylinder wrench.
Brass for flammables

## Slide 39
These are NOT TOOLS for cylinder cap removal or connecting regulators

## Slide 40: Attaching the Regulator continued
Apply an approved leak solution to the connection to test for any leaks
The regulator may now be adjusted to the correct pressure.
Safety Moment

## Slide 41: Does the regulator match my gas?
Heat of recompression
Another example of opening the cylinder valve slowly.

## Slide 42
Does the regulator match my application?
Using an industrial regulator on a specialty gas cylinder. It’s like drinking clean water through a dirty straw.
How can you tell the difference?

## Slide 43
My CO2 cylinder is leaking.

## Slide 44: Washer or Seals
Use a new washer (seal) every time you change a cylinder if the cylinder valve requires a washer.

## Slide 45: Left Hand = Counter-clockwise Has a V-Groove on the nut

## Slide 46: Regulator Mounting
Right Way
Wrong Way
The regulator juggle.

## Slide 47: Do you require continual gas flow?

## Slide 48: Gas equipment life span
Depending on gas service, materials used in construction, regulator manufacturer & customer usage frequency the normal specialty gas regulator has a life expectancy of 7-10 years.
Some customers change out the regulators every 5-7 yrs., depending on what type of gas is used & many other factors.
Most inert gas regulators (580) should be good for 7 yrs.
Corrosive gas regulators (330,660 and 705) can have a life expectancy of 90 days to 12 months.
Pigtails in inert gas service should be replaced every 7 years.

## Slide 49
Tubing & Fittings
Don’t forget the Material Compatibility

## Slide 50
On line Material Safety Data Sheets
Safety moment

## Slide 51: AERO
Airgas
Emergency
Response
Organization

## Slide 52: Thank you for your time. Have a safe day.
To change image:
- View / Slides Master
- Click on the image center and delete it
- Insert new image and send it to back
Position it and crop it to fit the shape
Do not use multiple photos or photo collage